## Sheet1
Колонка_1 | Колонка_2 | Колонка_3 | Колонка_4
--Test Summary--
Full Path and File Name
Instrument Name | 7222 #1
Instrument Type | Consistometer
Instrument Serial Number | 113
Test Start Date | 09.07.2002
Test Start Time | 20:19:58
Test Stop Date | 09.07.2002
Test Stop Time | 23:44:10
Test Length | 3:24:12
Final Sample Time Offset | 3:24:00
Controller Schedule Name
Sampling Period | 60 seconds
Comments | MPa: 25(30min); T: 50C(30min), tap water, dry-blending, cement from Sovetsky, Old Diacel FL
Information | WellID
Information | Customer | Petro-Aliance
Information | CementClass | G
Information | CementMfr | Dyckerhoff
Information | Additives | 0.4%Diacel FL+0.6%D-19
Information | Density | 1940
Information | BHCT | 50
Information | BHST | 60
Information | FluidLoss
Information | JobType | Research
Calculated Value | Thickening_Time | N/A
Calculated Value | Setting Time | N/A
Initial Temperature (C) | 29 C
Final Temperature (C) | 48 C
Initial Pressure (MPa) | 2 MPa
Final Pressure (MPa) | 25 MPa
Initial Consistency (Bc) | 17 Bc
Final Consistency (Bc) | 63 Bc
--Data--
Time (HH:MM:SS) | Temperature (C) | Pressure (MPa) | Consistency (Bc)
0:00:00 | 29 | 2 | 17
0:01:00 | 29 | 3 | 16
0:02:00 | 29 | 3 | 16
0:03:00 | 29 | 4 | 16
0:04:00 | 29 | 5 | 19
0:05:00 | 29 | 6 | 16
0:06:00 | 30 | 7 | 15
0:07:00 | 31 | 8 | 15
0:08:00 | 31 | 9 | 15
0:09:00 | 32 | 9 | 14
0:10:00 | 33 | 10 | 14
0:11:00 | 34 | 10 | 14
0:12:00 | 35 | 11 | 14
0:13:00 | 36 | 12 | 14
0:14:00 | 37 | 13 | 16
0:15:00 | 38 | 14 | 19
0:16:00 | 39 | 14 | 19
0:17:00 | 40 | 15 | 22
0:18:00 | 40 | 16 | 17
0:19:00 | 41 | 17 | 22
0:20:00 | 42 | 17 | 23
0:21:00 | 43 | 18 | 19
0:22:00 | 43 | 19 | 18
0:23:00 | 44 | 20 | 14
0:24:00 | 44 | 21 | 16
0:25:00 | 45 | 21 | 16
0:26:00 | 46 | 22 | 15
0:27:00 | 46 | 23 | 21
0:28:00 | 47 | 24 | 18
0:29:00 | 47 | 25 | 23
0:30:00 | 48 | 25 | 17
0:31:00 | 49 | 25 | 17
0:32:00 | 49 | 25 | 16
0:33:00 | 50 | 25 | 15
0:34:00 | 50 | 25 | 15
0:35:00 | 50 | 25 | 16
0:36:00 | 51 | 25 | 15
0:37:00 | 51 | 25 | 16
0:38:00 | 50 | 25 | 16
0:39:00 | 50 | 25 | 15
0:40:00 | 50 | 25 | 15
0:41:00 | 50 | 25 | 15
0:42:00 | 49 | 25 | 17
0:43:00 | 49 | 25 | 23
0:44:00 | 49 | 25 | 23
0:45:00 | 49 | 25 | 25
0:46:00 | 49 | 25 | 26
0:47:00 | 49 | 25 | 20
0:48:00 | 48 | 25 | 20
0:49:00 | 48 | 26 | 17
0:50:00 | 48 | 26 | 20
0:51:00 | 48 | 26 | 20
0:52:00 | 48 | 26 | 19
0:53:00 | 48 | 26 | 20
0:54:00 | 48 | 26 | 21
0:55:00 | 48 | 26 | 21
0:56:00 | 48 | 26 | 19
0:57:00 | 48 | 26 | 22
0:58:00 | 48 | 26 | 22
0:59:00 | 48 | 26 | 22
1:00:00 | 49 | 26 | 20
1:01:00 | 49 | 26 | 24
1:02:00 | 49 | 26 | 22
1:03:00 | 49 | 26 | 20
1:04:00 | 49 | 26 | 19
1:05:00 | 49 | 26 | 17
1:06:00 | 49 | 26 | 24
1:07:00 | 49 | 26 | 21
1:08:00 | 49 | 26 | 20
1:09:00 | 49 | 26 | 20
1:10:00 | 49 | 26 | 20
1:11:00 | 49 | 26 | 24
1:12:00 | 49 | 26 | 18
1:13:00 | 49 | 26 | 21
1:14:00 | 49 | 26 | 18
1:15:00 | 49 | 26 | 20
1:16:00 | 49 | 26 | 20
1:17:00 | 49 | 26 | 18
1:18:00 | 49 | 26 | 16
1:19:00 | 49 | 26 | 15
1:20:00 | 48 | 26 | 16
1:21:00 | 49 | 26 | 18
1:22:00 | 49 | 26 | 19
1:23:00 | 49 | 26 | 20
1:24:00 | 48 | 26 | 22
1:25:00 | 48 | 26 | 22
1:26:00 | 49 | 26 | 22
1:27:00 | 49 | 26 | 22
1:28:00 | 48 | 26 | 19
1:29:00 | 49 | 26 | 18
1:30:00 | 49 | 26 | 18
1:31:00 | 49 | 26 | 21
1:32:00 | 49 | 26 | 19
1:33:00 | 49 | 26 | 19
1:34:00 | 48 | 26 | 23
1:35:00 | 49 | 26 | 19
1:36:00 | 49 | 26 | 20
1:37:00 | 49 | 26 | 17
1:38:00 | 48 | 26 | 16
1:39:00 | 49 | 26 | 17
1:40:00 | 49 | 26 | 17
1:41:00 | 49 | 26 | 17
1:42:00 | 49 | 26 | 19
1:43:00 | 49 | 26 | 22
1:44:00 | 49 | 26 | 29
1:45:00 | 49 | 26 | 22
1:46:00 | 49 | 26 | 20
1:47:00 | 49 | 26 | 23
1:48:00 | 49 | 26 | 22
1:49:00 | 49 | 26 | 23
1:50:00 | 49 | 26 | 25
1:51:00 | 49 | 26 | 26
1:52:00 | 49 | 26 | 25
1:53:00 | 49 | 26 | 20
1:54:00 | 49 | 26 | 21
1:55:00 | 49 | 26 | 21
1:56:00 | 49 | 26 | 23
1:57:00 | 49 | 26 | 20
1:58:00 | 49 | 26 | 21
1:59:00 | 49 | 26 | 20
2:00:00 | 49 | 26 | 19
2:01:00 | 49 | 26 | 20
2:02:00 | 49 | 26 | 20
2:03:00 | 49 | 26 | 21
2:04:00 | 49 | 26 | 20
2:05:00 | 49 | 26 | 23
2:06:00 | 49 | 26 | 25
2:07:00 | 49 | 26 | 24
2:08:00 | 49 | 26 | 27
2:09:00 | 49 | 26 | 27
2:10:00 | 49 | 26 | 30
2:11:00 | 49 | 26 | 22
2:12:00 | 49 | 26 | 26
2:13:00 | 49 | 26 | 23
2:14:00 | 49 | 26 | 26
2:15:00 | 49 | 26 | 26
2:16:00 | 49 | 26 | 25
2:17:00 | 49 | 26 | 24
2:18:00 | 49 | 26 | 24
2:19:00 | 49 | 26 | 24
2:20:00 | 49 | 25 | 24
2:21:00 | 49 | 25 | 23
2:22:00 | 48 | 25 | 23
2:23:00 | 48 | 25 | 22
2:24:00 | 48 | 25 | 22
2:25:00 | 48 | 26 | 22
2:26:00 | 48 | 26 | 22
2:27:00 | 48 | 26 | 22
2:28:00 | 48 | 26 | 22
2:29:00 | 48 | 26 | 22
2:30:00 | 48 | 26 | 22
2:31:00 | 48 | 26 | 22
2:32:00 | 48 | 26 | 21
2:33:00 | 48 | 26 | 21
2:34:00 | 48 | 26 | 21
2:35:00 | 49 | 26 | 22
2:36:00 | 49 | 26 | 22
2:37:00 | 49 | 26 | 23
2:38:00 | 49 | 26 | 23
2:39:00 | 49 | 26 | 23
2:40:00 | 49 | 26 | 24
2:41:00 | 49 | 25 | 38
2:42:00 | 49 | 25 | 25
2:43:00 | 49 | 25 | 25
2:44:00 | 48 | 25 | 27
2:45:00 | 48 | 25 | 25
2:46:00 | 48 | 25 | 25
2:47:00 | 48 | 26 | 36
2:48:00 | 48 | 26 | 33
2:49:00 | 48 | 26 | 37
2:50:00 | 48 | 26 | 37
2:51:00 | 48 | 26 | 36
2:52:00 | 48 | 26 | 35
2:53:00 | 48 | 26 | 32
2:54:00 | 48 | 26 | 33
2:55:00 | 48 | 26 | 31
2:56:00 | 49 | 26 | 32
2:57:00 | 49 | 26 | 31
2:58:00 | 49 | 26 | 30
2:59:00 | 49 | 25 | 32
3:00:00 | 49 | 25 | 31
3:01:00 | 49 | 25 | 36
3:02:00 | 49 | 25 | 32
3:03:00 | 48 | 25 | 35
3:04:00 | 48 | 25 | 35
3:05:00 | 48 | 25 | 33
3:06:00 | 48 | 25 | 32
3:07:00 | 48 | 25 | 34
3:08:00 | 48 | 26 | 33
3:09:00 | 48 | 26 | 33
3:10:00 | 48 | 26 | 59
3:11:00 | 48 | 26 | 58
3:12:00 | 48 | 26 | 51
3:13:00 | 48 | 26 | 41
3:14:00 | 48 | 26 | 47
3:15:00 | 49 | 26 | 53
3:16:00 | 49 | 26 | 58
3:17:00 | 49 | 26 | 64
3:18:00 | 49 | 25 | 64
3:19:00 | 49 | 25 | 51
3:20:00 | 49 | 25 | 72
3:21:00 | 48 | 25 | 59
3:22:00 | 48 | 25 | 60
3:23:00 | 48 | 25 | 60
3:24:00 | 48 | 25 | 63